## Main List
SPARE PART LIST FOR AMS FOR ACE-ASIA FIELD CAMPAIGN
Jose Jimenez | 2001-01-23
COMPONENT | Supplier | Cost | Spares on hand | Spares ordered | Can Aerodyne Fedex them to Japan? | Failure probability | Failure history | $$ for Higher Priority Extra Spares | Pieces for which we would have NO spare | Notes
Parts w/ adequate level of spares
Electron Multiplier | ETP | 900 | 4 |  | Maybe | sure (~1) | 1 per month
Filaments | Balzers | 600 | 2 pairs | 3 pairs | Maybe | high (~1) | 1 per month
Chopper Servo | Hi-Tech | 50 | 5 |  | Maybe | high | 1 per 2 weeks
Chopper Motor (old) |  | 50 | 4 |  | Maybe | high | 1 per 2 weeks
Chooper Motor (new) | Micromo | 300 | 1 | 1 | Maybe | low | never
Oven | Heatwave | 400 |  | 1 | Maybe | low | never |  |  | Also ordered Moly Mesh and ribbon, need to spot weld
Turbo pump | Varian | 3800 | 1 |  | YES (1) | medium | 3 of prototype AMS, very hard duty |  |  | Can rebuild for $1000 in Japan
ConvecTorr | Varian | 300 | 2 |  | Maybe | low | 2 in AMS history
EM HV supply | EMCO | 300 | 1 |  | Maybe | low | never |  |  | Spare is 3 kV
Elec. Components in Circuits | Aerodyne | low | Assortment |  | YES | high | a few times |  |  | R, C, & transistors
Anal. Output Board | Nat. Inst. | 800 |  | 1 w/ Jay | Maybe | low | just started using it
Parts for which we could get spares if $$ is available. But Failure probability is low for most of them
Turbo pump controller | Varian | 800 | 1 |  | Maybe | low | 1 in AMS history | 800 |  | Would be safer to have an extra one
Fast DAQ Board | Nat. Inst. | 3000 |  |  | Maybe | low | never | 3000
Slow DAQ Board | Nat. Inst. | 500 |  |  | Maybe | low | 1 in AMS history | 500
High-vacuum gage | Varian | 800 |  |  | NO - Don't have it | medium | 2 failed, probably due to defect in gage | 800 |  | Has a hot filament, could be affected by vibration
Serial Board | Nat. Inst. | 800 |  |  | NO - Don't have it | low | just started using it | 800
Elec. Circuit 1 | Aerodyne | 300 |  |  | Maybe - Slow | medium | a few times | 300
Elec. Circuit 2 | Aerodyne | 300 |  |  | Maybe - Slow | medium | a few times | 300
Elec. Box Power Supplies | Various | 1000 |  |  | Maybe | low | never | 1000
Backing Pump | KNF | 2500 |  |  | Maybe | low | never | 2500 |  | J. Jayne Testing MD1, MD-60 DC from Vacubrand
Parts for which I was not planning to get spares. Some because of very low failure probability, some because they are very expensive
BNC Connector boxes | Nat. Inst. | 500 |  |  | Maybe | very low | never |  | 500
NI Cable | Nat. Inst. | 100 |  |  | Maybe | very low | never |  | 100
Computer | Dig.Dist. | 3000 |  |  | Maybe | medium | never |  | 3000
UPS | APC | 100 |  |  | NO - Don't have it | low | just started using it |  | 100 | Will not use if in-flight power is very stable
Quadrupole | Balzers | 20000 |  |  | NO - Maybe | low | never |  | 20000 | Could upgrade to bigger quad rods if spending this $$
RF Power Supply | Balzers | 8000 |  |  | NO - Maybe | low | never |  | 8000 | Could upgrade to bigger quad rods if spending this $$
Ionizer Power Supply | Balzers | 10000 |  |  | NO - Maybe | low | never |  | 10000 | Could upgrade to bigger quad rods if spending this $$
Ionizer Cable | Balzers | 400 |  |  | NO - Maybe | very low | never |  | 400
P-gage Controller | Varian | 2500 |  |  | Maybe | low | never |  | 2500
Flowmeter P-Transducer | Ashcroft | 400 |  |  | YES | very low | never |  | 400
Signal Amplifier | Adv.Res. | 1000 |  |  | YES - Maybe | very low | never |  | 1000
Custom Cables | Aerodyne | 500 |  |  | Maybe | low | never |  | 500 | for turbos, chopper, power
[7 empty] | TOTAL: | 10000 | 46500

## Essentials
PRIORITY LIST FOR AMS SPARE PARTS
Jose, Jan. 28, 2001
1 | High-vacuum gage | Varian | 800
2 | Elec. Circuit 1 | Aerodyne | 300
3 | Elec. Circuit 2 | Aerodyne | 300
4 | Turbo pump controller | Varian | 800
5 | Fast DAQ Board | Nat. Inst. | 3000
6 | Slow DAQ Board | Nat. Inst. | 500
7 | Serial Board | Nat. Inst. | 800
8 | Backing Pump | KNF | 2500
9 | Elec. Box Power Supplies | Various | 1000